ISSUES ON APPLICATIONS OF ANSI C12.1 AND C12.20

ANSI C12.1 and ANSI C12.20 have the following issues and allow the following test procedures to be performed which do not reflect the actual field conditions as intended:

- Not all components of meters are required to be included in the meters to be tested as a whole unit.

- Separate groups of meters can be performed separately under the same test.

- All tests can be performed in parallel so that only  total of one test can be performed on a meter.  This does not reflect the actual field conditions as any meter in service is being exposed to heat and cold cycles, humidity cycle, all surges, rain, sun, ect.

- A failure is about to occur and a repair can be made to prevent a failure to happen during a certification test.

- More than a minimum number of meters are allowed for testing, but test results are not required to be based on all meters tested.  Therefore, the test results can be submitted only on the meters that are tested successfully and less than the total meters tested as long as these meters are more than three as specified in ANSI C12.1.  Additionally, tests can be repeated indefinitely until three meters are tested successfully.

- There are no failure definitions of function performance even though functions are specified by manufacturer.  Even if any function specified by manufacturer is not being tested, but appears not being performed by the Meter during or after a certification, this could be counted as a non-failure (subject to interpretation).

- There are no failure definitions of firmware, software and hardware.  As a results, ANSI C12.1 allows a repair to be made to correct an undefined failure.

- There are no criteria of failure/success which allows indefinitely-repeated testing.