ANSI C12.20-1997

# AMERICAN NATIONAL STANDARD

FOR ELECTRICITY METERS

0.2 AND 0.5 ACCURACY CLASS

Accredited Standard Committee on Electricity Metering, C12

Accredited by

American National Standards Institute

Approved October 1997

American National Standards Institute, Inc.

American
National
Standard

Approval of an American National Standard requires verification by ANSI that the requirements for due process, consensus, and other criteria for approval have been met by the standards developer.

Consensus is established when, in the judgment of the ANSI Board of Standards Review, substantial agreement has been reached by directly and materially affected interests. Substantial agreement means much more than a simple majority, but not necessarily unanimity. Consensus requires that all views and objections be considered, and that a concerted effort be made toward their resolution.

The use of American National Standards is completely voluntary; their existence does not in any respect preclude anyone, whether he has approved the standards or not, from manufacturing, marketing, purchasing, or using products, processes, or procedures not conforming to the standards.

The American National Standards Institute does not develop standards and will in no circumstances give an interpretation of any American National Standard. Moreover, no person shall have the right or authority to issue an interpretation of an American National Standard in the name of the American National Standards Institute. Requests for interpretations should be addressed to the secretariat or sponsor whose name appears on the title page of this standard.

CAUTION NOTICE:  This American National Standard may be revised or withdrawn at any time. The procedures of the American National Standards Institute require that action be taken periodically to reaffirm, revise, or withdraw this standard. Purchasers of American National Standards may receive current information on all standards by calling or writing the American National Standards Institute.

Published by

National Electrical Manufacturers Association

1300 N. 17th Street, Rosslyn, Virginia  22209

© Copyright 1997 by the National Electrical Manufacturers Association. All rights including translation into other languages, reserved under the Universal Copyright Convention, the Berne Convention for the

Protection of Literary and Artistic Works, and the International and Pan American Copyright Conventions.

No part of this publication may be reproduced in any

form, in an electronic retrieval system or otherwise,

without prior written permission of the publisher.

Printed in the United States of America

TABLE OF CONTENTS

Foreword	iii

1	Scope	1

2	Definitions	1

3	References	1

4	Requirements	1

5	Acceptable performance	3

6	Requirements applicable to detachable electricity meters for 
	0.2 and 0.5 accuracy classes	21

7	Requirements applicable to bottom-connected 0.2 and 0.5 accuracy classes	21

8	Procedure for assigning form designations	21

Foreword	(This foreword is not part of ANSI C12.20-1997.)

This American National Standard provides acceptable performance criteria for electricity meters are established. Detachable socket, type "S", and bottom-connected, type "A", as well as any other arrangements agreed upon between the manufacturer and the user are included. Accuracy class designations, current class designations, voltage and frequency ratings, test current values, service connection arrangements, pertinent dimensions, form designations, and environmental tests are covered.

Suggestions for improvement to this standard are welcome. They should be sent to:

National Electrical Manufacturers Association

Vice President of Engineering

1300 North 17th Street

Suite 1847

Rosslyn, VA  22209

This standard was processed and approved for submittal to ANSI by Accredited Standards Committee for Electricity Metering, C12. At the time the committee approved this standard, the C12 Committee had the following members:

Brad Johnson, Chairman

Khaled Masri, Secretary

Organization Represented	Name of Representative

Drew, T.		ABB Power T&D Company

Beverly, W.		APTECH, Inc.

Jannelli, R.		Siemens PT&D

Johnson, B.		TU Electric

Malich, S.		Tampa Electric

McEvoy, J.		Florida Power & Light

Nguyen, D.Y.		Pacific Gas & Electric

Pananen, L.		Pacificorp

Powers, G.		Public Service Electric & Gas

Purc, M.		Schlumberger

Smith, C.		Pacificorp

Stevens, R.		Schlumberger

Vahlstrom, T.		Pacific Gas & Electric Co.

Buckley, B.		Itron

Buckner, E.		Mississippi Power

Germer, W.		GE

McDonald, K.		Georgia Power Company

Martin, J.		Scientific Columbus

Malemezian, E.		Florida Power & Light

Williams, D.		AVO International

for Electricity Meters 0.2 and 0.5 Accuracy Class

1	Scope

This standard establishes acceptable performance criteria for specification and performance requirements for 0.2 and 0.5 Accuracy Class, Electricity Meters meeting Blondel's Theorem. This standard does not apply to meters manufactured before the date of publication of this document.

2	Definitions

See Section 2 of ANSI C12.1-1995.

Note:	Definitions that are unique to this standard will be added.

3	References

Note:	The Reference Section will be completed at the end.

4	Requirements

4.1	Mounting

Mounting arrangements may include detachable socket type "S", bottom-connected type "A" or any other arrangement agreed upon between the manufacturer and user.

4.2	Voltage and frequency

Typical voltage and frequency ratings are 120, 240, 277 and 480 Volts and frequency rating of 60 Hz.

4.3	Current classes and test amperes

The current classes and test amperes shall be as listed in Table 1.

Table 1 – Current classes and test amperes

| Current Class | Test Amperes |
| --- | --- |
| 10 | 2.5 A |
| 20 | 2.5 A |
| 100 | 15 A |
| 200 | 30 A |
| 320 | 50 A |

NOTE: Other values of test amperes may be used as recommended by the manufacturer.

4.4	Form designations

The form designations for 0.2 and 0.5 accuracy classes, electricity meters are listed below.  Refer to ANSI C12.10-19XX for internal wiring diagrams.

1S

3S		3A

5S		5A

9S		9A

10S		10A

12S		12A

13S		13A

16S		16A

17S

25S

26S

35S		35A

45S		45A

48A

56S

4.5	Displays

4.5.1	Viewing requirements

Must be able to read all displayed elements under the following conditions with the cover in place:

- Vertical viewing angle; ±15o from the center of the display.

- Horizontal viewing angle; ±10o from the center of the display.

- Viewing conditions; Not closer than 3 feet from the face of the meter.

- Viewing conditions; Outdoor and Daylight.

4.5.2	Size of digits

The numeric digits for visual reading of billing quantities excluding identification codes shall not be less than 0.300 inch in height.

4.5.3	Direction of energy flow

A means to indicate direction of energy flow shall be provided.

4.6	Provisions of testing energy

Meters shall provide a test output proportional to energy.

4.7	Nameplates

Nameplates shall have the information as specified in ANSI C12.10-199X with the following additional items.

4.7.1	Meter nameplates shall contain the following information:

(1)	Accuracy Class

4.7.2	Permissible abreviations

CA  =	Accuracy Class

5	Acceptable performance

5.1	General

See paragraphs 4.1.2 through 4.6.2 in ANSI C12.1-1995.

5.1.1	Acceptable meters

In order to be acceptable, meters shall conform to the requirements specified in section 5.5, which are intended to determine their reliability and acceptable accuracy insofar as these qualities can be demonstrated by laboratory tests.

5.2	Types of meters

5.2.1	Meter type

Meters are considered to be of the same type if produced by the same manufacturer, and have similar measurement techniques and same uniform construction of parts determining these measuring techniques.

5.3	Conditions of test

5.3.1	Tests to be applied

The tests specified in 5.4 are to be conducted on the number of meters as specified in 4.4.2 of ANSI C12.1-1995.  Each individual meter is not required to be subjected to all tests.  Meters for special services or modifications to a type that have already been subjected to these tests, may be exempted from certain tests.  Replacement meters can be used for those meters that are found to fail some tests, be defective, or be damaged during test.

5.4	Performance requirements

5.4.1	Test conditions

The meter shall be mounted on a support that is free from vibration.  All alternating-current tests shall be conducted on a circuit supplied by a sine-wave source with a distortion factor not greater than 3%.  Where the meter has more than one voltage and current circuit, it shall be tested with the voltage circuits effectively in parallel and the appropriate current circuit(s) energized effectively in series, unless otherwise specified.

All tests shall be made at 23oC ±3oC, rated voltage ±3%, rated frequency ±1 Hz, test amperes ±3%, and unity power factor ±2°, unless otherwise indicated in specific tests.  The meter shall be stabilized at ambient temperature before performance tests are made.

5.4.2	Accuracy tests

5.4.2.1	Test no. 1:  No load

See 4.7.2.1 of ANSI C12.1-1995

5.4.2.2	Test no. 2:  Starting load

The meter shall operate continuously with a load current as specified in Table 2. The lowest rated voltage value should be used for wide voltage ranging meters.

Table 2 – Starting load test

|  | Accuracy Classes (Current in Amperes) |  |  |
| --- | --- | --- | --- |
| Current Class | 0.5 | 0.2 |
| 10 | 0.01 | 0.010 |
| 20 | 0.01 | 0.010 |
| 100 | 0.05 | 0.05 |
| 200 | 0.10 | 0.10 |
| 320 | 0.16 | 0.16 |

5.4.2.3	Test no. 3:  Load performance

The performance of the meter shall not deviate from the reference registration by an amount exceeding that specified in Table 3.

5.4.2.4	Test no. 4:  Effect of variation of power factor

Each element of a multi-element meter shall be tested as a single-element meter, except that all voltage circuits shall be in parallel.

5.4.2.4.1  Single-element meters

The effect of variation of power factor upon performance of the meter shall not exceed that specified in Table 4.

Table 3 – Load performance test

|  | Current Class |  |  |  |  |  |  |  | Maximum Deviation in Percent from Reference Performance |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | (Current In Amperes) |  |  |  |  |  | (Accuracy Class) |  |
| Condition | 10 | 20 | 100 | 200 | 320 |  | 0.5 | 0.2 |
| (1) | 0.15 | 0.15 | 1 | 2 | 3 |  | ±1.0 | ±0.4 |
| (2) | 0.25 | 0.25 | 1.5 | 3 | 5 |  | ±0.5 | ±0.2 |
| (3) | 0.50 | 0.5 | 3 | 6 | 10 |  | ±0.5 | ±0.2 |
| (4) | 1.5 | 1.5 | 10 | 20 | 30 |  | ±0.5 | ±0.2 |
| (5) | 2.5 | 2.5 | 15 | 30 | 50 |  | Reference | Reference |
| (6) | -- | 5 | 30 | 60 | 75 |  | ±0.5 | ±0.2 |
| (7) | 5 | 10 | 50 | 100 | 100 |  | ±0.5 | ±0.2 |
| (8) | 7.5 | 15 | 75 | 150 | 150 |  | ±0.5 | ±0.2 |
| (9) | -- | 18 | 90 | 180 | 250 |  | ±0.5 | ±0.2 |
| (10) | 10 | -- | 100 | 200 | 300 |  | ±0.5 | ±0.2 |
| (11) | -- | 20 | -- | -- | 320 |  | ±0.5 | ±0.2 |

Table 4 – Effect of variation of power factor for single-element meters

|  | Current Class |  |  |  |  |  | Factor |  | Maximum Deviation in Percent from Reference Performance |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | (Current in Amperes) |  |  |  |  | Power |  | (Accuracy Class) |  |  |
| Condition | 10 | 20 | 100 | 200 | 320 | Factor |  | 0.5 |  | 0.2 |
| Reference performance for Condition (1) Condition (1) | 0.25 0.5 | 0.5 1 | 1.5 3 | 3 6 | 5 10 | 1.0 0.5 lag |  | Reference ± 1.0 |  | Reference ± 0.5 |
| Reference performance for Condition (2) Condition (2) | 5 5 | 10 10 | 50 50 | 100 100 | 150 150 | 1.0 0.5 lag |  | Reference ± 0.6 |  | Reference ± 0.3 |
| Reference performance for Condition (3) Condition (3) | 10 10 | 20 20 | 100 100 | 200 200 | 320 320 | 1.0 0.5 lag |  | Reference ± 0.6 |  | Reference ± 0.3 |

5.4.2.4.2  Two-element meters

The effect of variation of power factor upon the performance of the meter shall not exceed that specified in Table 5.

5.4.2.4.3  Three-element meters

The effect of variation of power factor upon the performance of the meter shall not exceed that specified in Table 6.

5.4.2.5 Test no 5:  Effect of variation of voltage

The effect of variation of voltage upon the performance of the meter shall not exceed that specified in Table 7.

Table 5 – Effect of variation of power factor for two-element meters

|  | Current Class |  |  |  |  |  | Maximum Deviation in Percent from Reference Performance |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | (Current in Amperes) |  |  |  |  | Power | (Accuracy Class) |  |
| Condition | 10 | 20 | 100 | 200 | 320 | Factor | 0.5 | 0.2 |
| Reference performance for conditions (1) and (2) Condition (1) Condition (2) | 0.5 0.5 1.0 | 0.5 0.5 1.0 | 3 3 6 | 6 6 12 | 10 10 20 | 1.0 0.866 lead 0.5 lag | Reference ± 1.0 ± 1.0 | Reference ± 0.5 ± 0.5 |
| Reference performance for Condition (3) Condition (3) | 2.5 2.5 | 2.5 2.5 | 15 15 | 30 30 | 50 50 | 1.0 0.866 lead | Reference ± 0.6 | Reference ± 0.3 |
| Reference performance for conditions (4) and (5) Condition (4) Condition (5) | 5 5 5 | 10 10 10 | 50 50 50 | 100 100 100 | 150 150 150 | 1.0 0.866 lead 0.5 lag | Reference ± 0.6 ± 0.6 | Reference ± 0.3 ± 0.3 |
| Reference performance for Conditions (6) and (7) Condition (6) Condition (7) | 10 10 10 | 20 20 20 | 100 100 100 | 200 200 200 | 320 320 320 | 1.0 0.866 lead 0.5 lag | Reference ± 0.6 ± 0.6 | Reference ± 0.3 ± 0.3 |

Table 6 – Effect of variation of power factor for three-element meters

|  | Current Class |  |  |  |  |  |  | Maximum Deviation in Percent from Reference Performance |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | (Current in Ampere) |  |  |  |  | Power |  | (Accuracy Class) |  |
| Condition | 10 | 20 | 100 | 200 | 320 |  | Factor | 0.5 | 0.2 |
| Reference performance for condition (1) Condition (1) | 0.5 1 | 0.5 1 | 3 6 | 6 12 | 10 20 |  | 1.0 0.5 lag | Reference ± 1.0 | Reference ± 0.5 |
| Reference performance for Condition (2) Condition (2) | 5 5 | 10 10 | 50 50 | 100 100 | 150 150 |  | 1.0 0.5 lag | Reference ± 0.6 | Reference ± 0.3 |
| Reference performance for condition (3) Condition (3) | 10 10 | 20 20 | 100 100 | 200 200 | 320 320 |  | 1.0 0.5 lag | Reference ± 0.6 | Reference ± 0.3 |

Table 7 – Effect of variation of voltage

|  | Current Class |  |  |  |  |  |  |  | Maximum Deviation in Percent from Reference Performance |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | (Current in Amperes) |  |  |  |  | (Accuracy Class) |  |
| Condition | 10 | 10 | 100 | 200 | 320 | 0.5 | 0.2 |
| Reference performance 100% of calibration voltage for condition (1) and (2) | 0.2 | 0.25 | 1.5 | 3 | 5 | Reference | Reference |
| Condition (1) 90% of calibration voltage | 0.25 | 0.25 | 1.5 | 3 | 5 | ±0.2 | ±0.1 |
| Condition (2) 110% of calibration voltage | 0.25 | 0.25 | 1.5 | 3 | 5 | ± 0.2 | ± 0.1 |
| Reference performance 100% of calibration voltage for conditions (3) and (4) | 2.5 | 2.5 | 15 | 30 | 50 | Reference | Reference |
| Condition (3) 90% of calibration voltage | 2.5 | 2.5 | 15 | 30 | 50 | ±0.2 | ±0.1 |
| Condition (4) 110% of calibration voltage | 2.5 | 2.5 | 15 | 30 | 50 | ±0.2 | ±0.1 |

5.4.2.6	Test no. 6:  Effect of variation of frequency

The effect of variation of frequency upon the registration of a meter carrying constant load shall not exceed that specified in Table 8.

5.4.2.7	Test no. 7:  Equality of current circuits

5.4.2.7.1  Multi-element meters

The change produced in the performance of a multi-element meter by using only one current circuit as compared with that when all current circuits are used shall not exceed that specified in Table 9.

Table 8 – Effects of variation of frequency

|  | Current Class |  |  |  |  |  | Percent |  | Max. Deviation in % from Reference Performance |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | (Current in Amperes) |  |  |  |  |  | Rated |  | (Accuracy Class) |  |
| Condition | 10 | 20 | 100 | 200 | 320 | Frequency |  | 0.5 |  | 0.2 |
| Reference performance for conditions (1) & (2) | 0.25 | 0.25 | 1.5 | 3 | 5 | 100 |  | Reference |  | Reference |
| Condition (1) | 0.25 | 0.25 | 1.5 | 3 | 5 | 98 |  | ± 0.2 |  | ± 0.1 |
| Condition (2) | 0.25 | 0.25 | 1.5 | 3 | 5 | 102 |  | ± 0.2 |  | ± 0.1 |
| Reference performance for conditions (3) & (4) | 2.5 | 2.5 | 15 | 30 | 50 | 100 |  | Reference |  | Reference |
| Condition (3) | 2.5 | 2.5 | 15 | 30 | 50 | 98 |  | ± 0.2 |  | ± 0.1 |
| Condition (4) | 2.5 | 2.5 | 15 | 30 | 50 | 102 |  | ± 0.2 |  | ± 0.1 |

Table 9 – Equality of current circuits for multi-element meters

|  |  | Current Class |  |  |  |  | Max. Deviation in % from Reference Performance |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Connections of | (Current in Amperes) |  |  |  |  | (Accuracy Class) |  |
| Condition | Current Circuits | 10 | 20 | 100 | 200 | 320 | 0.5 | 0.2 |
| Reference performance for conditions (5), (6), (7), (8), etc. Condition (5) Condition (6) Condition (7), (8), etc. | All Circuits Circuit A only Circuit B only Circuits C, D, etc. | 0.25 0.25N* 0.25N* 0.25N* | 0.25 0.25N* 0.25N* 0.25N* | 1.5 1.5N* 1.5N* 1.5N* | 3 3N* 3N* 3N* | 5 5N* 5N* 5N* | Reference ±0.7 ±0.7 ±0.7 | Reference ±0.3 ±0.3 ±0.3 |
| Reference performance for conditions (9), (10), (11), (12), etc. Condition (9) Condition (10) Condition (11), (12), etc. | All Circuits Circuit A only Circuit B only Circuits C, D, etc. | 2.5 2.5 2.5 2.5 | 2.5 2.5 2.5 2.5 | 15 15 15 15 | 30 30 30 30 | 50 50 50 50 | Reference ±0.7 ±0.7 ±0.7 | Reference ±0.3 ±0.3 ±0.3 |

*N represents the number of elements in the meter

5.4.2.8	Test no. 8:  Internal meter losses

Refer to 4.7.2.8 of ANSI C12.1-1995.

5.4.2.9	Test no. 9:  Temperature rise

Refer to 4.7.2.9 of ANSI C12.1-1995.

5.4.2.10  Test no. 10:  Effect of internal heating

The test shall be made with the meter mounted in a conventional manner on a suitably rated meter mounting and wired with not less than 4 ft of conductor (8 ft jumper between terminals) of a size adequate for the load range of the meter.  Openings around the conductor, and any other openings, shall be closed with suitable material to prevent drafts.

The effect of internal heating upon the performance of a meter shall not exceed that specified in Table 10.

The reference test is to be made immediately upon energizing the meter.

Table 10a – Effect of internal heating

|  |  |  | Maximum Deviation in Percent from Reference Performance |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Current Class |  |  | Accuracy Classes |  |
| Condition | 10 | 20 |  | 0.5 | 0.2 |
| Reference performance for conditions (1), (2), and (7) | 10 | 20 |  | Reference | Reference |
| Reference performance for conditions (3) and (5) | 0.25 | 0.25 |  | Reference | Reference |
| Reference performance for conditions (4) and (6) | 2.5 | 2.5 |  | Reference | Reference |
| Condition (1) One-half hour after application of load | 10 | 20 |  | ±0.2 | ±0.2 |
| Condition (2) One hour after application of load | 10 | 20 |  | ±0.3 | ±0.2 |
| Condition (3) Immediately following test for condition (2) | 0.25 | 0.25 |  | ±0.3 | ±0.2 |
| Condition (4) Immediately following test for condition (2) | 2.5 | 2.5 |  | ±0.3 | ±0.2 |
| Condition (5) Two hours after test for condition (4) with meter at no load current during the two-hour interval | 0.25 | 0.25 |  | ±0.3 | ±0.2 |
| Condition (6) Immediately following test for condition (5) | 2.5 | 2.5 |  | ±0.2 | ±0.2 |
| Condition (7) Immediately following test for condition (6) | 10 | 20 |  | ±0.2 | ±0.2 |

Table 10b – Effect of internal heating for current classes 100, 200, and 300

|  | Current Class |  |  | Maximum Deviation in Percent from Reference Performance |  |
| --- | --- | --- | --- | --- | --- |
|  | (Current in Amperes) |  |  | Accuracy Classes |  |
| Condition | 100 | 200 | 300 | 0.5 | 0.2 |
| Reference performance for conditions (1), (2), and (7) | 100 | 200 | 320 | Reference | Reference |
| Reference performance for conditions (3) and (5) | 1.5 | 3 | 5 | Reference | Reference |
| Reference performance for conditions (4) and (6) | 15 | 30 | 50 | Reference | Reference |
| Condition (1) One-half hour after application of load | 100 | 200 | 320 | ±0.4 | ±0.4 |
| Condition (2) One hour after application of load | 100 | 200 | 320 | ±0.5 | ±0.4 |
| Condition (3) Immediately following test for condition (2) | 1.5 | 3 | 5 | ±0.5 | ±0.4 |
| Condition (4) Immediately following test for condition (2) | 15 | 30 | 50 | ±0.5 | ±0.4 |
| Condition (5) Two hours after test for condition (4) with meter at no load current during the two-hour interval | 1.5 | 3 | 5 | ±0.5 | ±0.4 |
| Condition (6) Immediately following test for condition (5) | 15 | 30 | 50 | ±0.4 | ±0.4 |
| Condition (7) Immediately following test for condition (6) | 100 | 200 | 320 | ±0.4 | ±0.4 |

5.4.2.11  Test no. 11:  Effect of polyphase loading

At full load and light load, with unity power factor, the difference between registration of ABC rotation, CBA rotation, and series load shall not exceed more than 0.6% for 0.5 accuracy class and 0.3% for 0.2 accuracy class.

5.4.3	External influences

5.4.3.1 Performance verification

Time, program and register readings shall be stored in the meter and/or register to be used as reference for tests 5.5.3.2 through 5.5.3.19.  After each test, proper operation of the meter and/or register shall be verified by time, program, and register reading check, and the meter shall show no damage and shall operate correctly in accordance with the requirement of the standard.  Any change in energy and power quantities shall be limited to ± 1 least significant digit displayed.

5.4.3.2	Test no. 12:  Insulation

Refer to 4.7.3.1 of ANSI C12.1-1995.

5.4.3.3	Test no. 13:  Voltage interruptions test

Refer to 4.7.3.2 of ANSI C12.1-1995.

5.4.3.4	Test no. 14:  Effect of high voltage line surges

Refer to 4.7.3.3 of ANSI C12.1-1995.

5.4.3.5	Test no. 15:  Effect of external magnetic field

Refer to 4.7.3.4 of ANSI C12.1-1995.

5.4.3.6  Test No. 16:  Effect of Variation of Ambient Temperature

The test shall be applied to three meters.  The meters shall be placed in a space having a temperature of 23oC ± 5oC and allowed to stand for not less than 2 hours with the voltage circuits of the meters energized.  Reference performance at each of the loads specified in Table 11 shall be obtained after operating the meters for 1 hour at the specified load.  The meters shall then be operated and tested at each of the following conditions:

- Conditions (1) through (6).  These tests shall be made with the meter placed in a space having a temperature of 50oC ± 5oC.  After energizing the voltage circuits of the meters for a period of not less than 2 hours, the appropriate test currents at the power factors listed for conditions (1) through (6) of Table 11 shall be sequentially applied to the meters.  Each condition shall be maintained for a period of at least 1 hour before tests are made to determine the deviation from reference performance.

- Conditions (7) through (12).  Repeat conditions (1) through (6), respectively, except that meters shall be placed in a space having a temperature of -20oC ± 5oC.  The effect of variation of temperature upon the performance of the meters shall not exceed that specified in Table 11.

Nominally available voltage and current can be used for the duration of this test when the metering device is not undergoing the specified accuracy test.

Table 11 – Effect of variation of ambient temperature

|  | Current Class |  |  |  |  |  |  |  |  |  |  |  | Max. Deviation in % from Reference Performance at Nominal Temp. Difference* |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | (Current in Amperes) |  |  |  |  | Power | Ambient |  | (Accuracy Class) |  |  |  |
| Condition | 10 | 20 | 100 | 2110 | 320 | Factor | Temperature |  | 0.5 |  | 0.2 |  |
| Reference performance for conditions (1) and (7) | 0.25 | 0.25 | 1.5 | 3 | 5 | 1.0 | 23oC±5oC |  | Reference |  | Reference |  |
| Reference performance for conditions (2) and (8) | 2.5 | 2.5 | 15 | 30 | 50 | 1.0 | 23oC±5oC |  | Reference |  | Reference |  |
| Reference performance for conditions (3) and (9) | 5 | 10 | 50 | 100 | 150 | 1.0 | 23oC±5oC |  | Reference |  | Reference |  |
| Reference performance for conditions (4) and (10) | 0.5 | 0.5 | 3 | 6 | 10 | 0.5 lag | 23oC±5oC |  | Reference |  | Reference |  |
| Reference performance for conditions (5) and (11) | 2.5 | 2.5 | 15 | 30 | 50 | 0.5 lag | 23oC±5oC |  | Reference |  | Reference |  |
| Reference performance for conditions (6) and (12) | 5 | 10 | 50 | 100 | 150 | 0.5 lag | 23oC±5oC |  | Reference |  | Reference |  |
| Condition (1) | 0.25 | 0.25 | 1.5 | 3 | 5 | 1.0 | 50oC±5oC |  | ±0.8 |  | ±0.3 |  |
| Condition (2) | 2.5 | 2.5 | 15 | 30 | 50 | 1.0 | 50oC±5oC |  | ±0.8 |  | ±0.3 |  |
| Condition (3) | 5 | 10 | 50 | 100 | 150 | 1.0 | 50oC±5oC |  | ±0.8 |  | ±0.3 |  |
| Condition (4) | 0.5 | 0.5 | 3 | 6 | 10 | 0.5 lag | 50oC±5oC |  | ±1.4 |  | ±0.5 |  |
| Condition (5) | 2.5 | 2.5 | 15 | 30 | 50 | 0.5 lag | 50oC±5oC |  | ±1.4 |  | ±0.5 |  |
| Condition (6) | 5 | 10 | 50 | 100 | 150 | 0.5 lag | 50oC±5oC |  | ±1.4 |  | ±0.5 |  |
| Condition (7) | 0.25 | 0.25 | 1.5 | 3 | 5 | 1.0 | -20oC±5oC |  | ±1.3 |  | ±0.5 |  |
| Condition (8) | 2.5 | 2.5 | 15 | 30 | 50 | 1.0 | -20oC±5oC |  | ±1.3 |  | ±0.5 |  |
| Condition (9) | 5 | 10 | 50 | 100 | 150 | 1.0 | -20oC±5oC |  | ±1.3 |  | ±0.5 |  |
| Condition (10) | 0.5 | 0.5 | 3 | 6 | 10 | 0.5 lag | -20oC±5oC |  | ±2.1 |  | ±0.9 |  |
| Condition (11) | 2.5 | 2.5 | 15 | 30 | 50 | 0.5 lag | -20oC±5oC |  | ±2.1 |  | ±0.9 |  |
| Condition (12) | 5 | 10 | 50 | 100 | 150 | 0.5 lag | -20oC±5oC |  | ±2.1 |  | ±0.9 |  |

*When the actual temperature difference between two tests differs from the nominal temperature difference specified for the two tests, the deviation shall be adjusted proportionally.

Table 12 – Effect of temporary overloads on accuracy

|  | Current Class |  |  |  | Maximum Deviation in Percent from Reference Performance |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | (Current in Amperes) |  |  | (Accuracy Class) |  |
| Condition | 100 | 200 | 320 | 0.5 | 0.2 |
| Reference performance for Condition (1) | 15 | 30 | 50 | Reference | Reference |
| Reference performance for Condition (2) | 1.5 | 3 | 5 | Reference | Reference |
| Condition (1) | 15 | 30 | 50 | ±0.1 | ±0.1 |
| Condition (2) | 1.5 | 3 | 5 | ±0.1 | ±0.1 |

Table 13 – Effect of current surge in ground conductor

|  | Current Class |  |  |  | Maximum Deviation in Percent from Reference Performance |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | (Current in Amperes) |  |  | (Accuracy Class) |  |
| Condition | 100 | 200 | 320 | 0.5 | 0.2 |
| Reference performance | 15 | 30 | 50 | Reference | Reference |
| Condition (1) | 15 | 30 | 50 | ±0.1 | ±0.10 |

5.4.3.7	Test no. 17:  Effect of temporary overloads

5.4.3.7.1  Effect on accuracy

Self-contained meters shall be subjected to short-circuit current of 7000 A peak, 60 Hz, for not more than 6 cycles (0.1 second).  For this test, current circuits of the meter shall be connected series adding.

The effect of short-circuit current on the performance of a meter should not exceed that specified in Table 12.

In order to eliminate residual effect, it is essential that tests of condition (1) be conducted before tests of condition (2).

5.4.3.7.2  Effect on mechanical structure and insulation

The meter shall withstand, for a duration of 4 cycles (0.07 second), a 60 Hz symmetrical fault current as follows without damage to the mechanical structure or reduction in the insulation level:

Current Class 100:10 000 A rms

Current Class 200:12 000 A rms

Current Class 320:12 000 A rms

5.4.3.8	Test no. 18:  Effect of current surge in ground conductor

Three meters shall be subjected to one transient surge of 20 000 A (20/50 microsecond wave) of either polarity through a conductor placed vertically 1.5 inches behind the flat portion of the base of the meter, with a socket in place.  This test shall not apply for meters with terminal compartments (Type A) or for Class 10 or Class 20 meters. The effect of current surge upon the performance of a meter shall not exceed that specified in Table 13.

5.4.3.9	Test no. 19:  Effect of superimposed signals

Refer to 4.7.3.8 of ANSI C12.1-1995.

5.4.3.10  Test no. 20:  Effect of voltage variation-secondary time base

Refer to 4.7.3.9 of ANSI C12.1-1995.

5.4.3.11  Test no. 21:  Effect of variation of ambient temperature-secondary time base

Refer to 4.7.3.10 of ANSI C12.1-1995.

5.4.3.12  Test no. 22:  electrical fast transient/burst test

Refer to 4.7.3.11 of ANSI C12.1-1995. The meter shall meet the Fast Transient Surge Test requirements of IEC 1000-4-4 (1995), for Electrical Fast Transient/Burst. See Figures 1 and 2.

This test subjects the power inputs and I/O circuits of the meter to repetitive bursts of 5 nanosecond rise time, 50 nanoseconds duration.

The test shall be carried out according to IEC 801-4, under the following conditions:

- Test severity level:  4;

- Test voltage on the voltage and current circuit:  4kV;

- Test voltage on I/O signal, data and control lines:  2 kV;

- Duration of the test:  minimum 60 seconds

- Voltage and auxiliary circuits energized with reference voltage;

- Without any current in the current circuits and the current terminals should be open circuit;

This test shall be conducted on all metering devices containing solid state components excluding voltage indicators.

5.4.3.13 Test no. 23:  Effect of radio frequency interference

This test shall be conducted with all cables connected, with all options and accessories specified, in a configuration closely resembling typical field in-service connections. See Figures 3, 4, and 5 for screen room testing and Figure 6 for GTEM!™ testing.  This test shall be conducted on all metering devices containing solid-state components excluding voltage indicators. For approval testing, the manufacturer need only use one test procedure, screen room or GTEM. Utility verification should use the same procedure as the manufacturer

5.4.3.13.1  Basic radiation susceptibility test

The test sample shall be subjected to both vertical and horizontal polarized continuous wave signals over a frequency range of 200 KHz-1000 MHz with a field strength of  15 ± 5 V/m.  The test shall be performed with the antenna facing the most sensitive side of the meter.  The field may be generated by a linearly polarized antenna positioned vertically and again with the antenna positioned horizontally, or a circularly polarized antenna may be used to simultaneously provide both vertical and horizontal susceptibility testing over those frequency ranges where circular polarized antennas are available, or with a uniform field.  The test procedures shall be conducted with samples configured so that disturbances can be readily noted.  Optional (add-on) functions and circuitry shall be installed when this represents a normal configuration for the test sample.  Special internal wiring or wire routing of the test sample are prohibited.

The test fixture shall be composed of a minimum of metal or other EMI reflecting or absorbing material capable of shielding or otherwise distorting the field in the vicinity of the test sample.  If a uniform field strength is not available, the fixture must allow the test sample to be oriented in each of  10 orientations to allow both horizontal and vertical irradiation of the front, left side, right side, top, and bottom of the test sample.

Below 50 MHz, the signal must be 90% amplitude modulated with a 1.0 KHz sine wave.  A continuous wave signal is used above 50 MHz.  Beginning at 200 KHz, the scan rate will double the frequency no faster than every 200 seconds (.005 octaves/second) through 1000 MHz.

One test sample is used to determine the orientation which provides the greatest sensitivity.

With a field strength set to 15 ± 5 V/m, test the test samples over a frequency range in the most sensitive orientation.

In each orientation above, the fixture must place the test sample in a test volume in which the field strength remains within acceptable test limits.

Use Figures 3 through 5 for typical shielded room testing and Figure 6 for GTEM testing.

When using GTEM testing, power and I/O leads should be only as long as needed to make connection from the GTEM shielded I/O connectors to the test sample.  Leads shall not exceed 5 meters in length.  Excess leads shall be coiled in a 0.4 meter diameter coil and supported  0.02 to 0.1 meter off the floor and other metal surface of the GTEM.

5.4.3.14  Test no. 24:  Radio frequency conducted and radiated emission test

Refer to 4.7.3.13 of C12.1-1995.

Figure 1 – Electrical fast transient/burst test #22

Figure 2 – Electrical fast transient/burst test #22 use of capacitive clamp to test data and I/O lines

Figure 3 – Typical test layout for radiated susceptibility,

radiated emissions, and conducted emissions

Figure 4 – Typical wiring detail self contained meters for radiated
susceptibility, radiated emissions, and conducted emissions

Figure 5 – Typical wiring detail transformer rated meters for radiated 
susceptibility, radiated emissions, and conducted emissions

Notes :
1 - Uniformity of the test field must be demonstrated to insure field does not fall below minimum field requirements.
2 - All power and cabling exiting the enclosure must be suitable filtered to prevent RF leakage.
3 - Orientation of the test meter is shown for front, horizontal polarization.
4 - Field intensity is controlled by septum voltage (direct measurement) or by field probes located in the cell.
5 - Field strength should be determined from the central axis of the meter : Voltage on septum divided by 
     septum height equals field strength.

Figure 6 – GTEM test

5.4.3.15  Test no. 25:  Effect of electrostatic discharge (ESD)

Refer to 4.7.3.14 of ANSI C12.1-1995.

5.4.3.16  Test no. 26:  Effect of storage temperature

Refer to 4.7.3.15 of ANSI C12.1-1995.

5.4.3.17  Test no. 27:  Effect of operating temperature

Refer to 4.7.3.16 of ANSI C12.1-1995.

5.4.3.18  Test no. 28:  Effect of relative humidity

Refer to 4.7.3.17 of ANSI C12.1-1995.

5.4.3.19  Test no. 29:  Mechanical shock

Refer to 4.7.3.18 of ANSI C12.1-1995.

5.4.3.20  Test no. 30:  Transportation drop

Refer to 4.7.3.19 of ANSI C12.1-1995.

5.4.3.21  Test no. 31:  Vibration

Refer to 4.7.3.20 of ANSI C12.1-1995.

5.4.3.22  Test no. 32:  Transportation vibration

Refer to 4.7.3.21 of ANSI C12.1-1995.

5.4.3.23  Test no. 33:  Weather simulation test

Refer to 4.7.3.22 of ANSI C12.1-1995.

5.4.3.24  Test no. 34:  Salt-spray test

Refer to 4.7.3.23 of ANSI C12.1-1995.

5.4.3.25  Test no. 35:  Raintightness

Refer to 4.7.3.24 of ANSI C12.1-1995.

6	Requirements applicable to detachable electricity meters for 0.2 and 0.5 accuracy classes

For construction requirements, refer to ANSI C12.10-19XX Section 4.

7	Requirements applicable to bottom-connected 0.2 and 0.5 accuracy classes

For construction requirements, refer to ANSI C12.10-19XX Section 5.

8	Procedure for assigning form designations

Refer to ANSI C12.10-19XX Section 7.