Draft Standards Testing

II.E.9. 	Requirements on Watt-hour Standards

The following are the requirements for maintenance and calibration of Portable watt-hour standards which are used for testing the accuracy of electric meters in the field or in the shop:

II.E.9.1 Out-of-calibration: A portable standard that is found to have accuracy greater than +/-0.1% of the basic reference standard shall be considered out-of-calibration.

II.E.9.2 Basic Reference Standard: A standard that is certified according to NIST requirements and used as an accuracy reference in the laboratory or meter shop.

II.E.9.3 Portable Standard: A standard that is used to certify the accuracy of meters in the field.

II.E.9.4 Rountine accuracy check requirements: Each test function of a portable standard shall be checked and compared with the basic reference standard annually.  If a portable standard used in testing meters has resulted in three consecutive meters being out of the CPUC required accuracy limits, this portable standard shall be certified against the basic reference standard in the laboratory.

Once a test standard is checked and found out-of-calibration, it shall be re-certified according to NIST requirements and manufacturer's recommendations, and meters that were tested by this portable standard since the last routine certification shall be re-tested with an accurately-calibrated portable watt-hour standard.